## Slide 1: Modelling Cardiac Pacemaker Software: Towards Methodology for Complex Critical Systems
Mike Poppleton
Abdolbaghi Rezazadeh
Colin Snook

## Slide 2: Roadmap
Background & motivation
Requirements ?
Electrocardiology basics
Requirements domain modelling – state diagrams
Solution domain modelling – different state diagrams
  Event-B brief recap
  Modelling & refinement … towards method
Further work

## Slide 3: Cardiac pacemaker: a definition
The device monitors and regulates a patient’s heart rate.
The device detects and provides therapy for bradycardia conditions.
The device provides programmable, single- and dual-chamber, rate-adaptive pacing, both permanent and temporary.
In adaptive rate modes, an accelerometer is used to measure physical activity, resulting in a sensor indicated rate for pacing the heart.

## Slide 4: Background & motivation
One of the Grand Challenge case studies

## Slide 5: Motivation
“In 1996 (in the USA), 10% of all medical device recalls were caused by software-related issues. In June of 2006, software errors in medical devices made up 21% of recalls.”
Jiang et al
In one study, 40% of cardiac complications in pacemaker-fitted patients were caused by the device

## Slide 6: The specification (Boston Scientific): lots of
this (definitions):

## Slide 7: The specification: and lots of
this (configuration data):
… ie very low-level, and little
behaviour is specified!

## Slide 8
[Cardiac Pacing and ICDs, Kenneth A Ellenbogen and Mark A Wood]

## Slide 9
[Cardiac Pacing and ICDs, Kenneth A Ellenbogen and Mark A Wood]

## Slide 10: Related work etc

## Slide 11: Motivation: related work
A number of teams have modelled the pacemaker, some down to code/ install on test devices
Not much evidence of refinement
Not much evidence of abstraction; cue taken from concrete style of requirement doc ?
No evidence of systematic reuse
At least 2 teams(Méry/Singh, Jiang et al) have modelled the heart (SUC) in conjunction with the pacemaker (CTL)
  Méry/Singh have used Event-B, refinement and proof (but in a concrete style)
  Jiang et al have used hybrid models and simulation

## Slide 12: Motivation: ours
Opportunities for systematic reuse (23 operational modes with common behaviours)
In search of more explicit repeatable modelling methodology
State diagram modelling
  gives elegant process models of requirements
  which map simply to guide modelling and refinement in Event-B
  gives a syntactic vehicle for structured reuse
Exploration of synchronisation modelling
Investigation of modelling styles for time – starting with Cansell/Mery

## Slide 13: To date we have:
Modelled 5 modes
  initially in CSP, then in a requirements-domain state diagram notation
  interpreted these in the solution domain of Event-B
  then refined down to introduction of explicit time intervals
  Model/verify(proof, model-check)/remodel cycle
Concerns:
  no (more) requirements analysis (than anyone else)

## Slide 16
<- Also known as AEI

## Slide 18: Behaviour: timing-driven sensing and pacing(DDD)

## Slide 19
RefA
RefA
RefA
RefA
MonV
MonA
MonA
MonA
PaceA
RefV
PaceV
PaceV
PaceV
RefV
MonV
RefV
MonV
MonV
RefV
SenseA
SenseA
SenseA
SenseV

## Slide 20: Requirements domain model – AAI

## Slide 21: Requirements domain models – AAI, VVI
AAI
VVI
NB - reuse: AAI -> VVI is simple refactoring by renaming

## Slide 22: Requirements domain model - VDD
NB: Synchronisation scheme between channels
NB – reuse: Each channel is an extended subgraph of AAI modulo renaming

## Slide 23: Event-B Overview
Two main elements: Contexts & Machines
Context: static model data and its logical properties
Machine: dynamic model variables (state), invariants on variables, state-updating events (behaviour)
Refinement: formal design transformation – layer in requirements, add architectural/implementation structure, (algorithmic) behaviour

## Slide 24: Handling Complexity

## Slide 25: Solution domain AAI : Abstract state machine

## Slide 26: Requirements to solution domain AAI : Abstract state machine

## Slide 27: AAI: Event-B Abstract Spec.

## Slide 28: AAI – State machine rep. of the first refinement

## Slide 29: AAI - First refinement

## Slide 30: AAI - First refinement(cont.)

## Slide 31: AAI – State machine rep. of the timing refinement

## Slide 32: Requirements -> solution domain mapping

## Slide 33
Abstract
Refined

## Slide 34: AAI: Second refinement – Introducing Timing properties

## Slide 35: VDD Mode - Reuse
A Channel
V Channel
Generic Models

## Slide 36: VDD Mode – synchronisation – refinement 1
A Channel
V Channel

## Slide 37
Requirements elicitation: ProB animation and model-checking of first refinement showed excessive event enablement, i.e. :
While A int-signal being processed, V pacing and int-signal processing is suspended
A sensed A int-signal must be (i) reported to V channel for synchronisation/suspension, and (ii) processed in a timely manner. (iii) Completion of processing must be reported to V channel for synchronisation/resumption
And vice-versa for V int-signal or pace
These are requirements on
signal flags pace_signal, intrinsic_signal_A, intrinsic_signal_V
synch flags a2vSignal, v2a_signal
VDD Mode – synchronisation – refinement 1

## Slide 39: VDD Mode – synchronisation – refinement 1a
A Channel
V Channel
Statically encode these requirements as safety invariants in 2 refinement steps
Strip out signals
Stop channels from being simultaneously active:

## Slide 40: VDD Mode – synchronisation – refinement 1b
A Channel
V Channel
Replace signals
Replace direct cross-chamber references with signal properties of each state :

## Slide 41: VDD Mode – Event-B Rep.

## Slide 42: VDD Mode – Event-B Rep.

## Slide 43: Conclusion
We have:
modelled AAI, VVI, DVI, VDD, DDD modes
  in a flat requirements-domain state diagram notation
  mapped this to a refinement chain of state machines in the solution domain (of Event-B)
explored loose synchronisation modelling (for concurrent code)
sketched a method pattern of dynamic -> static verification loop
identified statemachine-based reuse patterns
progressed statemachine design & animation tool support for Event-B
started investigation of adapted (from Cansell/Mery) modelling style for time

## Slide 44: Further work
Continue:
Synchronisation/ implementation design
Timing elaboration
State machine plugin & animation development
Functional development: hysteresis, rate modulation
Systematize statemachine-based reuse, develop tool support (building on the feature composition plugin)
Heart modelling for simulation-based testing of pacemaker
Utility of process (eg CSP) modelling ?
Continuous modelling of sensing / pacing mechanics ?

## Slide 45
That’s it folks.